## Slide 1: MRAD 251: CT Theory 1
Outcome 4:
Identify quality control requirements and quality
assurance data for computed tomography (CT)

## Slide 2: Objectives
Distinguish between the principles of quality control and quality assurance
Describe the principles of quality control
Describe data collected for quality assurance
Identify different types of quality control tests
Explain the purpose of quality control tests

## Slide 3: Objective 4.1
Quality Assurance is a management program used to ensure excellence in healthcare
A systematic
QA focuses on preventing defects or deficiencies
Quality Control is part of the QA program that deals with techniques used in monitoring and maintaining the technical elements of the systems that affect the quality of the image
QC deals with
Focuses on detecting deficiencies and defects

## Slide 4: Objective 4.2
QC is the part of a QA program that tests the performance of the CT scanner against an industry or machine standard.
The goal of QC is to ensure all acquired images are diagnostic with minimal dose to the patient
  QC is an important part of radiation protection

## Slide 5: Benefits of QC tests
QC ensures a scanners performance and accuracy
Image quality issues can be solved faster by evaluating the QC results
QC tests are also valuable for service engineers as the problems are often identified in the results

## Slide 6: QC Schedule
For a QC program to run smoothly:
QC tests need to be performed on a regular basis
  May be required to be performed: daily, weekly, monthly, quarterly, semi-annually, or yearly
Results need to be recorded & maintained
  Log all test results
Results must be interpreted promptly
  Comparing QC results to both the specifications of the scanner and previous test results
    Results can demonstrate if the scanners performance is degrading
      Analysis of results determines required response

## Slide 7: Objective 4.3
Quality Assurance (QA) is a program used by management to maintain optimal diagnostic image quality with minimum hazard and distress to patients. Can include: quality control, preventative maintenance, administrative processes, training of personnel, assessment of efficacy and quality of services, corrective processes. The primary goal of a radiology quality assurance program is to ensure the consistent provision of prompt and accurate diagnosis of patients. It is important that all technologists understand and contribute to QA; ensuring a culture of patient safety, ensuring optimal image quality, and effective and efficient workflow.
- CAMRT competency profile

## Slide 8: Objective 4.4
Acceptance testing:
(Usually done by a physicist)
  Ensures the scanner was delivered as ordered
    Aka. meets manufacturers specifications & radiation protection
regulations
  Tests the stability, integrity & safety of the scanner
Acceptance tests include:
•
•
Slice thickness verification
CT number linearity
Spatial and contrast resolution Noise
Dose output

## Slide 9: Steps in QC (cont’d)
Routine Performance Evaluation:
(Can be performed by the technologist and/or the physicist)
  Can be completed periodically
    Example: daily, weekly, monthly, and yearly
  Monitors the components of a CT scanner that affect image quality and dose
Error Correction:
  If the values of the QC tests are outside the acceptance or tolerance limits, the scanner needs to be serviced

## Slide 10: Performing QC Tests
Authorities such as the American College of Radiology (ACR), the International Atomic Energy Agency (IAEA), and Health Canada determine which QC procedures are to be performed by the technologist, and which will be performed by medical physicists
The manufacturer of the scanner may also recommend evaluative procedures

## Slide 11: QC Testing Frequency
General Rule for testing frequency:
Less important, more time consuming tests, are usually performed less frequently
  Monthly
The most complex tests are usually performed semi-annually, or annually
  These tests are usually performed by a physicist and may require the
use of QC phantoms
QC tests and their acceptance limits are also outlined in the Canadian SC 35

## Slide 12: Three Mandates of QC
Must be performed on a regular basis.
Schedule of QC testing is a compromise between pros & cons
  Scanner downtime, regulatory requirements & expected benefits
Prompt interpretation of measurements
  Must recognize when the CT scanner is not operating within acceptable limits
Consistent & accurate record keeping
  Results should be maintained for as long as the CT scanner is active

## Slide 13: Discerning Acceptable Limits
Perhaps more important than the actual value of the measured variable, is a change in the variable .
Two performance standards for QC test results:
  Acceptable
    Indicates that performance must be within certain tolerances
    If not, the equipment should not be used
  Achievable
    Indicates the level of performance that should be attained under favorable
circumstances
    The level at which a facility should work, if feasible

## Slide 14: QC Equipment
Phantoms are regularly used in QC testing.
Types of phantoms:
  Image performance
  Geometric
  Dosimetry
  Quantitative

## Slide 15: Techniques for QC Measurements
Depends on the type of CT scanner and the test being performed.
Suggestions when choosing technique combinations:
  Select one or two representative techniques
  Use techniques that match a frequently used clinical technique
    Example:
  Technique should remain the same for a specific test from day to day

## Slide 16: Test Priority and Frequency
Test priority classifications:
Essential
Desirable
Test frequency considerations:
Replacement of major components, equipment service or adjustments
  Requires immediate testing
If used for accurately localizing tissue (biopsies & radiation therapy treatment)
  ↑frequency of appropriate tests
Complexity of tests
  ↓frequency of test completion

## Slide 17: Test Frequency
Daily QC tests (as outlined by the Canadian SC 35) include:
Equipment Warm Up
  Usually a predefined protocol built into the scanner as per the manufacturers specifications
Meter Operation
Equipment Conditions
Visual inspection of electronic display devices
  Ensures the integrity of radiation protection considerations are maintained
  Frequency depends on regulatory body
    ACR states monthly, IAEA states performance by physicist, CAMRT states daily technologist responsibility

## Slide 18: Daily QC Procedures
Equipment Warm-up:
There is a warmup period before the system performs optimally
  Systems that have no “downtime” still require warm-up every 24hours
    Must be consistent, occurring at the same time of day
May include the boot-up of the system and operating system
Warms up the x-ray tube
Calibrates the detectors

## Slide 19: Visual Inspection Checklist
A visual inspection may include the following:
Checking the integrity of cables & wires
Checking the WW & WL display
Checking for appropriate warning labels & service records
Room cleanliness/safety concerns
Checking to ensure all equipment is functioning
  Examples:
    Table height & position indicators
    Gantry angulation indicators
    Laser localization lights
    Intercom system

## Slide 20: Weekly QC Tests
CT Number Accuracy
Using a uniform water phantom:
  The mean CT number of water and the SD within a large ROI must remain within the established baseli2ne & acceptable limits of variation
    CT number for H O must be within the range of 0 (± 4HU)
CT Noise
Using a uniform water phantom:
  Noise is calculated by the variation of CT numbers from a mean value in a
defined area in the image of the phantom
    Its magnitude is equal to the SD of the CT number values within the ROI
SD within the ROI must be no greater than ±10% (or 0.2HU)

## Slide 21: Weekly QC Tests (cont’d)
Uniformity
Using a phantom of water or uniform density
  Uniformity is defined as the consistency of the CT numbers in an image
of a homogeneous material across the scan field
  Uniformity is calculated using the difference in the mean CT number in the center ROI from those at the periphery of the image
    No more than ± 2HU variation from the center ROI to any at the periphery
      If measuring water, than the variation should be no greater than ± 5HU

## Slide 22: Monthly QC Tests
CT Slice Thickness
Using a phantom containing a ramp, spiral or step wedge of known
sizes/thicknesses
  Evaluates the tomographic section thickness
    If “selected” slice thickness is ≥ 2mm:
      The “measured” slice thickness should not vary by more than ±1mm from the
selected slice thickness
    If “selected” slice thickness is between 2mm→1mm:
      The measured slice thickness must not vary by more than 50% of the selected slice thickness
    If slice thickness is ˂1mm:
      The measured slice thickness should not differ by more than ±0.5mm from the selected slice thickness

## Slide 23: Monthly QC Tests (cont’d)
CT Number Linearity
Using a phantom containing uniform objects of known materials with a wide range of CT numbers
Tests are completed at all clinically used voltage settings
  The measured CT numbers of the materials should be compared with the values provided by the phantom manufacturer & with previously measured values
The measured values must remain within the established limits for the CT
scanner
  CT numbers checked over the CT range of: -1000 → +1000
  Acceptable limits achieved when values “form a straight line”
    Plot of average CT numbers as a function of linear attenuation coefficients

## Slide 24: Quarterly QC Tests
CT Spatial Resolution
Evaluated by:
Determining MTF
  Performed by a physicist
    Measurements of the 50% point (on the high contrast wire test tool) and the 10% point of the MTF curve must be within 0.5 lp/cm or ±15% of the established baseline
Imaging a line pairs phantom
  Performed by a technologist
    Using a line-pair test pattern with varying spatial frequencies at a point where the
measured MTF drops to 5%; the limiting resolution should be 5 lp/cm
At minimum, spatial resolution should be equal to the manufacturers specifications for the scanner

## Slide 25: Quarterly QC Tests (cont’d)
CT Low Contrast Resolution
Using a phantom of objects of varying size
  Objects have small density differences
    ˂1% or 10HU contrast to the surrounding material
Value determined as the smallest sized object seen at a specified contrast level to the background when imaged under specific conditions
  Should be evaluated using the reconstruction algorithm routinely used, as well as
other clinically relevant algorithms
CT Patient Support Movement
Test is performed using an even weight load of no more than 135kg
  A distance of 30cm (maximum scan increment) and return to start position
must not exceed ±1mm

## Slide 26: Semi - Annual QC Tests
Laser Light Accuracy
Lasers are located inside & outside of the gantry
  Used for patient positioning and alignment
The alignment of the laser lights to the radiation must be ˂
±2mm
Isocenter measurements must be within the manufacturers recommended tolerances with a limit of ±5mm

## Slide 27: Semi-Annual QC Tests
Accuracy of Automatic Position of Tomographic Plane
Assessment of scout images
  Ensures accuracy of the Localizer with the digital display location
  Reference lines indicating tomographic sections must not differ from true position by anymore than ±2mm with the gantry in vertical position

## Slide 28: Annual QC Tests
CT Number Dependence of Phantom Position
To ensure HU accuracy for varied patient positions within the gantry
  CT number for water must not vary by more than ±5HU when the phantom moves through the gantry
Radiation Dose Profile
Ensures the collimation of the radiation beam does not exceed the prescribed scan width

## Slide 29: Objective 4.5
Purpose of QC tests
Maintain high performance standards for patients.
Regulatory bodies demand maintaining a standard, due to the potential harm to patients and the public from ionizing radiation

## Slide 30: Things to Take Away
Safety Code 35 is a publication by Health Canada.
Defines “Safety Procedures for the Installation, Use and Control of X-ray Equipment in Large Medical Radiological Facilities”
  SC 35 is a reference used by the CAMRT for their exam questions for MRT certification
For more info on QC and testing, access the online SC35 as a resource (link information found in textbook readings list on Edenbrook; direct hyperlink access can be found on Edenbrook “QA & QC”resources tab)