APPH 226 - Lab 14 – Tube Conversion – ANSWER KEY

Radiographic Rating Charts and Anode Cooling Curves

| Generator Type | Modification Factor |
| --- | --- |
| Single Phase | 1 |
| Three Phase | 1.4 |
| High Frequency | 1.4 |

Formula to calculate HU:

HU = kVp x mA x time x modification factor

Complete the following questions while referring to the “Radiographic Rating

Chart”

- What is the maximum kVp that can be used at 600 mA and 3 s? 65 kVp

- Assuming a high frequency generator, how many HU would be generated by the exposure in question 1?

HU = kVp x mA x time x modification factor

65 kVp x 600 mA x 3s x 1.4 = 163 800 HU

- What is the maximum kVp that can be used at 400 mA and 4 s? 87-89 kVp

- Are the following exposure technique selections possible with this x-ray tube?

a. 100 kVp, 600 mA, 0.5 s	Yes / No

| b. 70 kVp, 300 mA, 5 s | Yes / No |
| --- | --- |
| c. 85 kVp, 700 mA, 0.5 s | Yes / No |

MRT Program – APPH 226

| d. 60 kVp, 500 mA, 0.9 s | Yes / No |
| --- | --- |
| e. 150 kVp, 600mA, 0.005s | Yes / No |
| f. 115 kVp, 250 mA, 0.01 s | Yes / No |
| g. 130 kVp, 400 mA, 1 s | Yes / No |
| h. 95 kVp, 800 mA, 0.2 s | Yes / No |

- What is the maximum time that can be used at:

  - 120 kVp and 500 mA 0.35 s

  - 90 kVp and 600 mA 0.75 s

  - 75 kVp and 400 mA 6 s

  - 90 kVp and 250 mA 10s

  - 105 kVp and 600 mA 0.3 s

  - 85 kVp and 600 mA 1 s

  - 100 kVp and 700 mA 0.16 s

  - 60 kVp and 600 mA

4 s

MRT Program – APPH 226

Complete the following questions while referring to the “Anode Cooling Curve”

- What is the maximum anode storage capacity (HU) for this tube? 350 000 HU

- Assuming this is a high frequency generator, what is the maximum number of consecutive exposures that could be made, using 80 kVp and 20 mAs?

kVp x mA x time x modification factor = HU

80 x 20 x 1.4 = 2240 HU

  - 000 / 2240 = 156.25

156 exposures

- What is the anode cooling time needed if the anode is presently at 315 000 HU and the exposure that is needed is 100 kVp, 450 mA, 0.8 s assuming this is a three phase generator?

kVp x mA x time x modification factor = HU 100 kVp x 450 mA x 0.8s x 1.4 = 50 400 HU

  - 000 - 50 400 = 299 600 HU

According to chart , ~ 20 s needed to drop from 315 000 HU to 299 600 HU

- What is the anode cooling time needed if the anode is presently at 300000 HU and the exposure that is needed is 95 kVp, 600 mA, 0.4 s assuming this is a single phase generator?

kVp x mA x time x modification factor = HU 95 kVp x 600 mA x 0.4 s x 1 = 22800

  - 000 – 22800 =  327 200 HU .. currently at 300 000 HU, so no need to wait No cooling time required, as total HU will be under 350 000.

MRT Program – APPH 226

- What is the anode cooling time needed if the anode is presently at max and the exposure that is needed is 120 kVp, 1600 mA, 0.5 s assuming this is a high frequency generator?

kVp x mA x time x modification factor = HU 120kVp x 1600 mA x 0.5 s x 1.4 = 134 400

350 000 – 134 400 = 215 600

According to chart, will require ~ 1.5 min to drop from 350 000 HU to 215 600 HU

- Assuming this is a three phase generator, what is the maximum number of consecutive exposures that could be made, using 120 kVp, 1000 mA and 0.8s?

kVp x mA x time x modification factor = HU 120kVp x 1000 mA x 0.8 s x 1.4 = 134 400 HU

  - 000 / 134 400 =2.6

2 exposures

- How long will a technologist need to wait if the anode is presently at 250000 HU and the exposure that is needed is 60 kVp, 600 mA, 4 s assuming this is a three phase generator?

kVp x mA x time x modification factor = HU 60 kVp x 600 mA x 4 s x 1.4 = 201600 HU

  - 000 – 201 600 = 148 400 HU

According to chart, will require ~1.5 min to drop from 250 000 HU to 148

400HU

MRT Program – APPH 226