APPH 226 – Apparatus and Image Management

Lab 14 – Tube Conversion – Extra Questions

Complete the following questions while referring to the “Radiographic Rating Chart”

- What is the maximum kVp that can be used at 900 mA and 0.1 s?

- Assuming a high frequency generator, how many HU would be generated by the exposure in question 1?

- Are the following exposure technique selections possible with this x-ray tube?

  - 120 kVp, 500 mA, 0.5 s			Yes / No

  - 80 kVp, 600 mA, 1 s			             Yes / No

  - 85 kVp, 400 mA, 0.5 s			Yes / No

  - 70 kVp, 800 mA, 1s				Yes / No

  - 120 kVp, 400mA, 0.05s			Yes / No

- What is the maximum time that can be used at:

  - 110 kVp and 600 mA

  - 70 kVp and 900 mA

  - 70 kVp and 400 mA

  - 95 kVp and 600 mA

  - 60 kVp and 500 mA

Complete the following questions while referring to the “Anode Cooling Curve”

- Assuming this is a high frequency generator, what is the maximum number of consecutive exposures that could be made, using 70 kVp and 12 mAs?

- What is the anode cooling time needed if the anode is presently at 300 000 HU and the exposure that is needed is 120 kVp, 400 mA, 0.6 s assuming this is a three phase generator?

- What is the anode cooling time needed if the anode is presently at 325 000 HU and the exposure that is needed is 90 kVp, 800 mA, 0.8 s assuming this is a single phase generator?

- What is the anode cooling time needed if the anode is presently at max and the exposure that is needed is 110 kVp, 1200 mA, 0.75 s assuming this is a high frequency generator?