# Activity: Monitor Device Chart – ANSWER KEY

| Type | How it Works | Advantages | Disadvantages |
| --- | --- | --- | --- |
| Film Badges | - 3 parts: holder, metal filters & film packet - Filters allow measurement of approximate energy of radiation by appearance of an image - Can also determine direction similar to dental film | - Economical, permanent legal record of personnel exposure - Can be used for x-radiation, gamma radiation - Can discriminate types, energies of radiation - Durable - Can indicate whether scatter radiation, or single exposure from primary | - Doses less than 0.1 mSv usually not detected - Only records exposures in area where worn - High temp/humidity can fog film - Cannot determine exposure on same day of occurrence - Accuracy +/- 20% |
| Thermoluminescent Dosimeter (TLD) Most commonly seen | - Looks like film badge - Contains crystals of lithium fluoride (sensing material) - Radiation causes crystals to change physical properties, trapping electrons - When heated, emit energy in form of light | - Not affected by humidity, pressure or temp - Can be worn up to 3 months - Crystals can be reused - Crystals interact as human tissue does, more accurate dose measurement | - Readings can be lost if not carefully recorded - Taking readings destroys info in TLD (not a permanent record) - Records exposure only in the area where it is worn - Exposure cannot be determined on same day of occurrence |
| Optically Stimulated Luminescence (OSL) Most commonly seen | - Aluminum oxide detector - Read with laser light at different frequencies - Becomes luminescent in proportion to amount of radiation exposure received - Plot light intensity and crystal heating temp to give exposure amount | - Lightweight, durable - Not affected by heat, moisture or pressure - Gives accurate reading as low as 1 mrem - More sensitive than a TLD, good for low radiation environments - Can be used up to 1 year | - Exposure cannot be determined on same day of occurrence |
| Pocket Ionization Chambers | - Looks like a fountain pen - Two electrodes: 1 +, 1 – - Quartz fiber can function as the indicator - Air around central electrode ionized and discharges mechanism in proportion to the radiation received | - Small, compact and easy to use - The most sensitive type of dosimeter - Provides immediate readout, good for high exposures | - Expensive - Uncommon in DI (used in high exposure areas, i.e. cardiac cath.) - If not read every day or calibrated to zero, can give inaccurate reading - Can discharge due to mechanical shock, giving false reading - No permanent legal record |