## Slide 1: RADP 215: Radiation Protection
Building Design and Protective Barriers

## Slide 2: Department Design and Safety
Main priority is to protect personnel and the public from radiation levels above the current exposure limits.
Routine checks ensure that construction and shielding is adequate.
There are many considerations when constructing a radiology department
Notes: Location, Adjoining rooms, controlled versus uncontrolled, use etc

## Slide 3: Location Within the Institution
Main floor, ground level
Upper floor, second or above
Depends on the infrastructure in the existing or new department
Adjoining rooms
Must consider patients that will be seen at the facility

## Slide 4: Department Design and Safety
All radiographic rooms must have a detailed floor plan that indicates:
Location of the control panel in the room
  Dimensions, shape and exact location in regard to other work areas
Location of the equipment within the room
Range of movement of the tube

## Slide 5: Types of Radiation
Types of radiation in an exam room:

## Slide 6: Primary Radiation
  Direct radiation (useful beam)
  Travels without deflection toward wall, door or viewing window
  Most intensive and most difficult to stop—requires the heaviest shielding
  Calculated using:
  Total primary = mR/mAs x mA x time

## Slide 7: Scatter Radiation
  Main source of dose to a technologist
  Caused when primary beam interacts with an object, causing it to change direction
  Patient is the single most important scattering object in the room
  Scatter can be regarded as a new source with a different intensity—approximately 1/1000 of the useful beam @ 1 meter
  Calculated using:
  mR @ patient x 0.1% @ 1m

## Slide 8: Leakage Radiation
  Radiation emitted from the tube housing in all directions other than the port
  Tube leakage must never exceed 1 mGy/hour or 0.1 R/hour or 100 mR/hour @1 meter
  Levels are typically much less than limit
  Calculated using:
  100 mR/hour @ 1 m

## Slide 9: Protective Barriers
Material to be used for the barrier affects the design of the building/department
Two common materials used to form barriers:

## Slide 10: Calculating Shielding
Comprehensive shielding calculations should only performed by individuals with current knowledge of structural shielding design and the acceptable methods of performing these calculations.
Appendix III in SC 35 for more information

## Slide 11: Calculating Barrier Thickness
All of the following factors must be considered in the barrier calculations:
  Distance (D) from the radiation source
  Controlled vs. uncontrolled areas
  Occupancy (T) on other side of the barrier
  Barrier type
  Use (U) of the unit or time the tube is energized
  Workload (W) of the unit or number of exams/cases done in a given time frame

## Slide 12: Distance (D)
Thickness depends on:
  Distance between the barrier and radiation source
  Where the tube is situated when energized
    Upright detector (wall)
    Table detector (floor)
    X-table (wall)

## Slide 13: Controlled Areas
Radiographic examination rooms and work areas
Often restricted entrance
  Radiology personnel (monitored)
  Patients, for a short time
SC 35 maximum exposure is 20 mSv/yr for controlled areas

## Slide 14: Uncontrolled Areas
Rest of the world!
Occupied by non-monitored personnel or the general public
SC 35 maximum exposure is 1 mSv/yr for uncontrolled areas
SC 35 states that the radiation-sensitive areas (pediatrics) should use a control level of 0.30 mSv/yr

## Slide 15: Occupancy factor (T)
Occupancy factor (T)
  Takes into account the time that the space behind the barrier is occupied in a work week
SC 35 Section B Table 6
Amount of time that room is occupied
  1 = full occupancy
  1/2 and 1/5 = partial occupancy
  1/8 = limited occupancy
  1/20 = occasional occupancy
  1/40 = infrequent occupancy

## Slide 16: Barrier Type
Primary Barrier
  Designed to absorb the uninterrupted primary beam
Secondary Barrier
  Designed to absorb misdirected scatter and leakage photons
Notes: Need to be able to protect from the most intense radiation
Required whenever the primary beam is directed at a structure
Any structure that the useful beam can be directed towards is designated a primary protective barrier
Most common is the wall behind upright detector

## Slide 17: Primary Barriers
Need to be able to protect from the most intense radiation
Required whenever the primary beam is directed at a structure
Any structure that the useful beam can be directed towards is designated a primary protective barrier
Most common is the wall behind upright detector

## Slide 18: Primary Barrier Materials
Lead bonded to sheet rock or wood paneling
Concrete, cinder blocks or bricks may be used in place of lead
Department floors are made of concrete that is thick enough to meet the criteria for a primary barrier
Thickness of lead vs. concrete
Shielding must be constructed to form an unbroken barrier

## Slide 19: Secondary Barriers
Protects from secondary radiation
  Leakage
  Scatter
Barrier doesn’t need to be as thick as primary barrier
Lead may not be required
Other options include:
  Gypsum board
  Glass
  Lead-infused acrylic

## Slide 20: Barriers
Shielding must be constructed to form an unbroken barrier
Must have 10-15 mm overlap at the seams
Must extend 2.1 meters from the floor
Must be permanently attached to the floor
Lead should be adequately supported to prevent “creeping”
Can be infused into the gypsum

## Slide 21: Use Factor (U)
Fraction of workload when the beam is on and is directed towards a barrier
Most structures in the room will be struck by radiation at some time throughout the week when the beam is on
Use factor is used to select the “fractional contact time” of the radiation
Section D Table 7

## Slide 22: Workload (W)
Weekly operational use of the X-ray unit
Measured in mA-min/wk or mAs/wk
  Section B 1.2.3 Table 5 SC 35
The busier the room, the more shielding required
  Busy radiology room
    > 500 mA-min/wk
  Slower clinic
    < 100 mA-min/wk

## Slide 23: Workload Units Calculation Example
X-ray room is in operation 5 days/week. The average number of patients/day is 20 and the average number of films/patient is 3. The average exposure factors are 90 kVp, 300 mA and 0.1 sec. Find the weekly workload:

## Slide 24: kVp
Calculations are done with a kVp higher than what is actually used in the department
  Average of 100 kVp for general X-ray
Calculations do not account for the attenuation of the beam when intercepted by the patient and image receptor

## Slide 25: Over-Shielding
Design must take into consideration all the previous factors
Plans should also take into consideration possible future changes:
  Advances in radiation technology
  ALARA rules becoming stricter
  For these reasons, medical imaging departments often over-estimate the required shielding in their calculations

## Slide 26: Design of X-ray Room
Must provide adequate working space for tech and patient transfer
Room door ______ be self-closing and _____ have a warning sign that states, “Unauthorized Entry Prohibited”
Unit should be placed so that when energized, no one can enter the room without the technologist knowing
Beam must only be directed toward adequately shielded primary barriers

## Slide 27: Design of X-ray Room
Wall detector barriers must have adequate primary shielding
When practical, primary and scatter radiation must be absorbed as close to the patient as possible
Control booth and window must provide protection from ________ radiation
Radiation has to be scattered at least ________before entering the control booth

## Slide 28: Purpose of Radiation Surveys
Evaluation done by an expert in radiation safety
Ensures the X-ray and auxiliary equipment is working within SC 35 standards.
Ensures proper equipment installation and provides maximum radiation protection for everyone

## Slide 29: Radiation Surveys
Evaluates shielding and protective devices
Equipment is assessed semi-annually or annually and submitted to the owner and/or responsible user
Reports must be presented in a clear, systemic way
Contains surveyor’s conclusions and recommendations

## Slide 30: Radiation Survey Report
Schematic of location, equipment, booth and surrounding areas
Equipment identification and type of tube support
Observation of electrical and mechanical operation
Actual or estimated total workload
Results of radiation readings, both in and outside of the controlled area

## Slide 31: Radiation Surveys
Assessment of X-ray, imaging or diagnostic performance, techniques, DRLs, total filtration, radiation protection and restraints
Estimate of potential exposures to personal and public
Investigates high OSL readings and potential of exposure to people
Review of QA program
Recommends improvements to above

## Slide 32: Radiation Protection Agencies
In order to obtain a registration certificate, radiation facilities and equipment need to be inspected and tested by an authorized radiation protection agency. Radiation protection agencies contract directly with the owners of the radiation equipment and are responsible for ensuring compliance with the Radiation Protection Act and Regulation.

## Slide 33: Radiation Emitting Devices Act
The Radiation Emitting Devices (RED) Act governs the sale, lease and import of certain radiation emitting devices used for medical and industrial purposes or by consumers.
The Act sets safety performance standards for the sale, lease, import, labelling, packaging, and advertising of radiation emitting devices to ensure that workers and the public are not placed at risk.

## Slide 34: RED Act
Manufacturers and importers are required to notify the Minister of Health if a device does not comply with the regulations or creates a risk to any person.
In addition, a series of voluntary safety codes that apply to radiation emitting devices in the workplace have been developed to provide guidance to employers.