Activity: Test 5 Automatic Exposure Control (AEC) Reproducibility Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the sample data for a completed AEC Reproducibility Test.

- Complete the calculations below.

- Answer the questions after the calculations.

## Evaluation of Data

| Exposure # | mA | mAs | Time (ms) | Exposure in mR |
| --- | --- | --- | --- | --- |
| 1 | 100 | 1.557 | 15.57 | 13.25 |
| 2 | 100 | 1.555 | 15.55 | 13.32 |
| 3 | 100 | 1.605 | 16.05 | 13.31 |
| 4 | 100 | 1.608 | 16.08 | 13.24 |
| 5 | 100 | 1.605 | 16.05 | 13.28 |
| 6 | 100 | 1.557 | 15.57 | 13.34 |

- Determine the percent of variation using the following formula:

% Variation = (mR max - mR min)   x 100

(mR max + mR min)

=        (              -                )     x 100

(               +             )

- Compare results to SC 35 and Papp, 2015, pp. 111–113.

- SC 35 states _________________________ should be used.  Papp states the reproducibility variance for an AEC device should be within _______.

## Critical Thinking Questions

- What is the main function of an AEC system?

- What is an advantage to using AEC systems in radiology departments?

- Why do we tend not to use AEC on distal extremities or even when certain pathologies are present?