# Activity: Test 14 Lead Apparel Testing Calculations/Evaluation

## Purpose

This activity will provide you with practice in viewing lead apparel to determine if it passes evaluation and meets the standards for use. This test is typically done fluoroscopically.

## Instructions

- Review the images below.

- The images would be evaluated to determine if the lead apparel is acceptable for use.

## Sample Evaluation

- On each image, check for areas of increased radiographic density which may indicate holes, cracks or wear. A description of the damaged areas needs to be documented.

- When determining the pass/fail of each item, take into consideration the size and location of the defects. A defect on the very edge of a lead apron may not require immediate replacement, whereas a crack through the middle (in the abdominal or gonadal region) would require the apron to be removed from usage.

- Record all results on either a data table (as used by EDENBROOK) or any method utilized by the individual department.

## Sample Chart

| Lead | # | kVp | mAs | Lead/mm | P\F | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| Apron |  | 80 | 5 |  |  |  |
| Apron |  | 80 | 5 |  |  |  |
| Apron |  | 80 | 5 |  |  |  |
| ½ Apron |  | 80 | 5 |  |  |  |
| Hang. Apron |  | 80 | 5 |  |  |  |
| Glove |  | 80 | 5 |  |  |  |
| Glove |  | 80 | 5 |  |  |  |
| Lead sheet |  | 76 | 5 |  |  |  |
| Lead sheet |  | 76 | 5 |  |  |  |
| Lead sheet |  | 76 | 5 |  |  |  |
| Lead sheet |  | 76 | 5 |  |  |  |