Activity: Test 6 Automatic Exposure Control (AEC) Field Balance/Uniformity Calculations

## Purpose

This activity will provide you with practice in evaluating the equipment by calculating data, using formulas and determining if results fall within acceptable limits.

## Instructions

- Review the data for a completed AEC Field Balance/Uniformity Test.

- Complete the calculations below.

- Answer the question after the calculations.

## Evaluation of Data

| Cell | mAs | Time (ms) | Exposure (mR) | % Diff (ms) | % Diff (mR) | Pass | Fail |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Centre | 3.01 | 30.10 | 18.14 |  |  |  |  |
| Left | 3.163 | 31.63 | 19.24 |  |  |  |  |
| Right | 3.110 | 31.10 | 18.95 |  |  |  |  |
| Centre + Left | 3.113 | 31.13 | 18.69 |  |  |  |  |
| Centre + Right | 3.011 | 30.11 | 18.49 |  |  |  |  |
| Left + Right | 3.113 | 31.13 | 19.00 |  |  |  |  |
| All | 3.063 | 30.63 | 18.73 |  |  |  |  |
| Mean |  |  |  |  |  |  |  |

- Determine the percent (%) difference for the ms and mR using the following formula:

% Difference = (Measured) – (Mean)   x 100

(Mean)

- Compare results to SC 35 and Papp, 2015, pp. 111–113.

- SC 35 states _________________________ should be used. The AHS standard for acceptance is < 20%.

Conclusion: Pass ______________      Fail ________________