# APPH 226 – Apparatus and Image Management

# Lab 6 - Grids & Exposure Technique

### Purpose

Objective: 5.3 Evaluate grids and their effect on image quality.

Correlate how grids influence technical factor selection.

### Instructions

Exposure #1

- Enter the patient information into the console, select L-spine.

- Select wall detector.

- Turn AEC off and set 80kVp @ 32mAs.

- Center the tube and detector using an SID of 110cm.

- Collimate to a 20 x 25cm field size.

- Position the body part phantom upright on the cart and as close to the detector as possible.

- Center over the midline of the body part phantom.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI readout: _________.

Exposure #2

- Maintain the same setup as Exposure #1.

- Move the body part phantom about 15-20 cm away from the detector.

- Center to match the centering used in Exposure #1.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI readout: _________.

Image Evaluation:

- Describe any differences between the 2 images.

- What type of technique was simulated in Exposure #2?

- Explain which exposure produced the most radiographic density based on the EI/EI_s.

Exposure #3

Grid Ratio: 12:1

- Select table detector.

- Turn AEC off and set 80kVp @ 32mAs.

- Center the tube and detector using an SID of 110cm.

- Collimate to a 20 x 25cm field size, center over the midline of the body part phantom.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI readout: _________.

Exposure #4

- Maintain the same setup as Exposure #3.

- Ask an instructor to remove the grid from the table detector.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Record the EI readout: _________.

Image Evaluation:

- Describe any differences between the 2 images.

- Explain which exposure produced the most radiographic density based on the EI/EI_s.

Exposure #5

Grid Ratio: ________

- Select free cassette.

- Set 60kVp @ 5mAs.

- Place Grid #4 on top of two glove boxes, over the table.

- Move the detector to ensure it will not be exposed.

- Center the tube to the grid with an SID of 110cm

- Place the dosimeter under one corner of the grid.

- Collimate the field to the edges of the grid ensuring the dosimeter is included.

- Remove the grid.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the results from the dosimeter: _________ mR.

Exposure #6

Grid Ratio: ________

- Maintain the same setup as Exposure #5.

- Place Grid #9 on top of the glove boxes, centered in the light field.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the results from the dosimeter: _________ mR.

Exposure #7

Ratio: ________

- Maintain the same setup as Exposure #5.

- Replace Grid #9 with Grid #3 on top of the glove boxes, centered in the light field.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the results from the dosimeter: _________ mR.

Exposure #8

Ratio: ________

- Maintain the same setup as Exposure #5.

- Replace Grid #3 with Grid #4 on top of the glove boxes, centered in the light field.

- Ensure everyone is out of the room & the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure

- Record the results from the dosimeter: _________ mR.

Questions:

- The grids used have ratios of non-grid, 8:1, 10:1 and 12:1.

Based on the lab results, fill in the grid ratios for exposure 6, 7 and 8. How did you determine which grid ratio would match for each exposure?

- How did the grid ratio affect the amount of exposure that reached the dosimeter?

- An exposure was made without the use of a grid. A repeat exposure is required with a grid. What must be changed to maintain the radiographic density between the two exposures?

- What is the main advantage of using a grid?

- Identify one disadvantage of using a grid.

- What type of grid would have the greatest effect on a radiographic image?

- Which type of grid would be the least beneficial to a patient?