### Purpose

Radiography Equipment: Visual Inspection

Objective: 1.2 Locate basic components of X-ray generating equipment.

Become familiar with the radiographic equipment located within the lab, and ensure safe operation of the x-ray units.

### Instructions

Complete as much of the visual checklist as possible (some sections will require an exposure prior to completion). **Remember verbal permission with direct supervision from an instructor is required BEFORE any exposure. **

# Annual Verification Checklist for Medical Imaging Equipment

This report is for: Acceptance	Compliance

Department: Room:

Generator Model: Manufacturer: Country:

Serial #: Equipment Type:

Tube Information

Description: Serial #:

Model:

Date Manufactured: Country Manufactured:

Inherent Filtration:		mm Al

# General Information and Facility Requirements

Yes	No	N/A

- There is a certificate for the room.

- If (1) is "yes", the certificate must be posted in or close to the room.

Location: (manager’s office)

- Radiationshieldingisadequateandmeetsregulatory standards.

4.X-raywarningsignsarepostedonallentrancedoorstothe room.

- Protectiveapparelisavailabletopatientsandindividualsin theradiationarea

- AQAprogramexistsforthedepartment. 7.Aqualitymaintenanceprogramexistsforthedepartment

## Labelling

Generator

Yes	No	N/A

# General Requirements

  - Name of the Manufacturer

  - Model Designation

  - Serial Number

  - Date of Manufacture

  - Country of Manufacture

X-ray Tube

Yes	No	N/A

- Name of the Manufacturer

- Model Designation

- Serial Number

- Date of Manufacture

- Country of Manufacture

- Minimum permanent inherent filtration

- Are the anode and cathode clearly marked?

## Warning Signs

Yes	No	N/A

The x-ray control panel must have a permanent, visible and legible sign warning that hazardous x-rays are emitted, and unauthorized use is prohibited.

## Markings

Yes	No	N/A

All controls, meters, lights and other indicators related to the equipment operation must be clearly labelled or marked as to their function.

## Mechanical Stability

Yes	No	N/A

- The x-ray tube must be securely fixed and correctly aligned within the tube housing.

- The x-ray assembly and support must maintain their required positions without excessive drift or vibration during operation.

## Indicators and Indicator Lights

Yes	No	N/A

- There must be distinctive and separate indicators on the control panel that indicate:

  - when the control panel is energized and the tube is ready to produce x-rays

  - when x-rays are being produced.

  - if AEC is provided when that mode of operation is selected

  - if AEC mode is not selected, or does not exist, the selected loading factors must be indicated to the operator before an irradiation

  - if equipment is battery operated, whether enough charge is present for proper operation of the equipment.

- If more than one x-ray tube is controlled by one control panel:

  - a visual indicator on or near each tube housing to show that the x-ray tube is connected and ready to be energized

  - a visual indicator on the control panel to indicate which of the x-ray tubes is connected and ready to be energized

- An audible signal to indicate the termination of an exposure.

## Indication of Loading Factors

Yes	No	N/A

- X-ray equipment having adjustable loading factors must have meters or other indicators that enable determination of the tube voltage, current and time, or combinations of these.

- The loading factors must be displayed before, during and after the exposure.

## Irradiation Control

Yes	No	N/A

There must be an irradiation switch, controlling timer or another mechanism to initiate and terminate x-ray production.

## Irradiation and Emergency Switches

Yes	No	N/A

- When the equipment is equipped with an exposure switch, it must require continuous pressure by the operator to emit x-rays.

- In the case of a foot switch, it must prevent the emission of any unintended x-rays when it is overturned.

- In the case of mobile equipment, it must allow the operator to stand at least 3m away from the x-ray source when the x-ray tube is energized.

- If the equipment moves around the patient by remotecontrol, there must be an emergency stop switch thatimmediately terminates both equipment movement and

exposure.

## Controlling Timer

Yes	No	N/A

When the equipment has a controlling timer, it must be constructed so that:

  - it can automatically terminate an exposure on completion of a preset irradiation time, or a preset current time product value, or on completion of a preset number of x-ray pulses.

  - it permits the operator to terminate the exposure at any time

  - it automatically resets itself to its original setting or to zero on termination of an irradiation.

  - when it is at zero, in the off position, or at an unmarked setting, irradiation cannot occur.

## Focal Spot Marking

Yes	No	N/A

The location of the focal spot must be clearly and accurately marked on the x-ray tube housing.

# Testing Frequency

This is performed annually or in conjunction with tube and generator testing. A less extensive inspection should be done by the technologist daily or when required by the department protocol.