Activity Answer: Test 15 Visual Inspection of Fluoroscopic Equipment Procedure

Critical Thinking Question Answer

- Why should a fluoroscopic system visual inspection occur every 6 months, while a general X-ray visual inspection occurs annually?

- More moving parts, locks, motors for power assist etc.

- Exposure switch (especially foot pedals) can malfunction or stick, increasing dose to patient and technologist, radiologist

- Lights/meters indicate when fluoroscopy is activated, both visual and auditory, if defective can increase dose to patient and staff

- More potential for things to break down