APPH 226 – Apparatus and Image Management

Lab 14 – Tube Conversion

Radiographic Rating Charts and Anode Cooling Curves

| Generator Type | Modification Factor |
| --- | --- |
| Single Phase | 1 |
| Three Phase | 1.4 |
| High Frequency | 1.4 |

Formula to calculate HU:

HU = kVp x mA x time x modification factor

Complete the following questions while referring to the “Radiographic Rating Chart”

- What is the maximum kVp that can be used at 600 mA and 3 s?

- Assuming a high frequency generator, how many HU would be generated by the exposure in question 1?

- What is the maximum kVp that can be used at 400 mA and 4 s?

- Are the following exposure technique selections possible with this x-ray tube?

  - 100 kVp, 600 mA, 0.5 s			Yes / No

  - 70 kVp, 300 mA, 5 s			Yes / No

  - 85 kVp, 700 mA, 0.5 s			Yes / No

  - 60 kVp, 500 mA, 0.9 s			Yes / No

  - 150 kVp, 600mA, 0.005s			Yes / No

  - 115 kVp, 250 mA, 0.01 s			Yes / No

  - 130 kVp, 400 mA, 1 s			Yes / No

  - 95 kVp, 800 mA, 0.2 s			Yes / No

- What is the maximum time that can be used at:

  - 120 kVp and 500 mA

  - 90 kVp and 600 mA

  - 75 kVp and 400 mA

  - 90 kVp and 250 mA

  - 105 kVp and 600 mA

  - 85 kVp and 600 mA

  - 100 kVp and 700 mA

  - 60 kVp and 600 mA

Complete the following questions while referring to the “Anode Cooling Curve”

- What is the maximum anode storage capacity (HU) for this tube?

- Assuming this is a high frequency generator, what is the maximum number of consecutive exposures that could be made, using 80 kVp and 20 mAs?

- What is the anode cooling time needed if the anode is presently at 315 000 HU and the exposure that is needed is 100 kVp, 450 mA, 0.8 s assuming this is a three phase generator?

- What is the anode cooling time needed if the anode is presently at 300000 HU and the exposure that is needed is 95 kVp, 600 mA, 0.4 s assuming this is a single phase generator?

- What is the anode cooling time needed if the anode is presently at max and the exposure that is needed is 120 kVp, 1600 mA, 0.5 s assuming this is a high frequency generator?

- Assuming this is a three phase generator, what is the maximum number of consecutive exposures that could be made, using 120 kVp, 1000 mA and 0.8s?

- How long will a technologist need to wait if the anode is presently at 250000 HU and the exposure that is needed is 60 kVp, 600 mA, 4 s assuming this is a three phase generator?