APPH 226 – Apparatus and Image Management

# Lab 5 - Collimation (9 penny test)

### Purpose

Demonstrate the accuracy of the collimator.

### Instructions

Diagram for the set up:

Exposure #1

- Enter the patient information into the console, and select a Finger exam.

- Select table detector.

- Center the tube and detector at an SID of 110cm.

- Collimate to an 18 x 24cm field size.

- On the table top, place 4 pennies on the inside edge of the light field.

- Place another 4 pennies on the outside edge of the light field, so that they are touching the pennies in step 5.

- Place the 9th penny in the light field on the anode side.

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Use the post-processing mask function to identify how many pennies were clipped in the exposure.

Exposure #2

- Maintain the same setup as Exposure #1.

- Collimate to a field size of 24 x 30cm.

- Move the pennies to the edge of the light field (identical to the previous setup).

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Use the post-processing mask function to visualize all the pennies on the monitor.

Exposure #3

- Maintain the same setup as Exposure #1.

- Collimate to a field size of 35 x 43cm.

- Move the pennies to the edge of the light field (identical to the previous setup).

- Ensure everyone is out of the room and the door is closed.

- Obtain verbal confirmation with direct supervision from an instructor.

- Take the exposure.

- Use the post-processing mask function to visualize all the pennies on the monitor.

Questions:

- Does the light field pass between the coins in each of the exposures?

18 x 24cm:

24 x 30cm:

35 x 43cm:

- Calculate 2% of an SID of 110cm.

- When would the collimation be considered unacceptable?

- What is the purpose of marking the anode side with a penny?

- Use the measuring tool on your console to determine if the collimators are within acceptable limits.

18 x 24cm:

24 x 30cm:

35 x 43cm: