## Slide 1: MRAD 251: CT Theory 1
Outcome 5 (Part 2):
Summarize radiation safety in computed tomography
(CT)

## Slide 2: Objective
5.3 Explain the use of dosimetry in CT

## Slide 3: Objective 5.3
Dosimetry can be defined as: Instrumentation & methods used to measure patient dose from a CT scanner.
Dose Measurements can be used for:
  Risk assessment
  Radiation protection guidelines
  Dose optimization

## Slide 4: Importance of Dosimetry
Dosimetry enables technologists the ability to:
Compare their doses to national averages
  Regulatory bodies have an anonymous database for comparison
Determines effectiveness of the implementation of their radiation protection methods
Inform the public & other personnel about CT doses
Perform dose measurements
  When medical physicist is unavailable

## Slide 5: Importance of Dosimetry Cont’d
Diagnostic Reference Levels (aka. DRL)
Purpose:
  Maintain standards, used to control amount of radiation exposure, for
patient imaging
Considerations:
  Maximum dose limits should not be used
    Will not regulate safe dose limits for patient exposures
      Dose limit values would be too high
  Doses for procedures fluctuate due to variances in:
    Equipment
    Facility protocols

## Slide 6: Dosimetry Concepts
In order to appreciate the data collected from dosimetry measurements technologists should understand:
Types of dosimeters used to measure CT doses
CT dosimetry phantoms
Dose descriptors specific to CT
  Including their units of measurement

## Slide 7: 1. Radiation Dosimeters
Dosimeters are devices used for measuring exposure to ionizing radiation.
Two main uses:
  Human radiation protection
    Personal dosimeters
  Measurements of dose in medical processes
    Measures radiation emitted from CT scans

## Slide 8: Types of Dosimeters
Types of Dosimeters:
Film
TLD
Specially designed Ionization Chambers

## Slide 9: Film Dosimeters
Film badge dosimeter:
A type of Personal Dosimeter
Has two parts:
  Film & Holder
Double emulsion technology advancement
  Enabled detection of low & high energy photons
Least accurate form of measuring CT dose
  Sensitivities to light, heat, & humidity
  One time use only

## Slide 10: TLD
Thermoluminescent Dosimeter:
A type of Personal Dosimeter
Active crystal component:
  Detects radiation exposure
  Emits light when heated
    Light energy is proportional to the amount of radiation absorbed by the crystal
Reusable
  Can be worn for up to 3 months
  once heated & measured

## Slide 11: Ionization Chambers
Used to measure radiation doses from CT scans.
Specially designed Ionization Chambers:
  used for CT dose measurements
    Current method used for measuring CTDI
  Easiest method of recording exposure
  Most accurate method of quantifying radiation exposure

## Slide 12: Pencil Ionization Chamber
Small air-filled container with thin walls that allow radiation to pass through.
X-rays collide with air molecules within the chamber
Some of these molecules are ionized
The ionized electrons are collected on a conducting wire or plate and measured as an electric charge
The collected charge is proportional to the amount of ionization
  Which is proportional to the amount of radiation
  Q
The charge is removed from the chamber and measured by an electrometer

## Slide 13: 2. Dosimetry Phantoms
Standardize dose measurements for various CT exams.
Phantoms mimic patient geometry
Phantom Characteristics:
•
  Made of acrylic
Contain holes
  For placement of the pencil ionization chamber
  Unused holes need to be “plugged”
Enables dose measurements at different locations
Come in two diameters with the same length:
  16cm
  32cm

## Slide 14: Phantom Considerations
The dose measured from the phantoms using a consistent technique:
Varies
  Dependent on location of measurement
    Due to: Partial shielding
Measures CTDI
  Index of CT radiation dose
•
  Uses estimates to calculate Helical scan doses

## Slide 15: 3. CT Dose Descriptors
Three primary dose measurements for CT:
Computed Tomography Dose Index (CTDI)
  Units of measure = Grays (Gy)
Dose Length Product (DLP)
  Units of measure = mGy/cm (milligrays per centimeter)
Effective Dose (ED)
  Units of measure = Sieverts (Sv)

## Slide 16: I. CTDI
Dose is usually calculated from multiple scans.
CT scans require multiple slices
•
CTDI measures the MSAD (aka. multiple scan average dose)
  Calculates the dose (of total exposure to radiation) of a single slice
    Accounts for scatter within each slice
•
Provides an average estimated measurement of the exposure per slice of tissue

## Slide 17: I. CTDI (cont’d)
CTDI is a standardized measurement of radiation dose.
Allows dose comparison between scanners
•
  Estimates are used to calculate Helical scan doses
    Due to limitations caused by the geometry of pencil ionization
chambers
DOES NOT account for differences in tissue densities within a patient

## Slide 18: CTDI Measurements
Dose must be measured at several locations.
Increases accuracy of average dose estimates
Difference between CTDI measurements:
CTDIFDA
CTDI100
CTDI
W
•CTDVIolume
= mean absorbed dose in the scanned
object volume; fixed slice width measurements
= measures a variety of slice widths
= calculates average dose in the x-y axis
= calculates average dose in the z-axis
To calculate the average dose of a scan slice in Helical CT
imaging:
Volume
CTDIW
Pitch
CTDI =

## Slide 19: Calculating Helical “slice dose”
A Helical CT Head scan was acquired on a 6-slice MDCT, with a Pitch of 1.2. Calculate the
CTDIVol.
Step 1: write out the formula
Step 2: fill in the formula with numerical values
CTDIVol = ?
CTDIW =
P =
Step 3: complete the math and record your answer
CTDIVol = Gy
CTDI | 
Calculates dose in the x-y axis to accommodate dose uniformity | 0.024Gy
Modification to accommodate smaller slice widths | 0.013Gy
Dose index as it relates to number of slices & slice width used | 0.022Gy

## Slide 20: II. DLP
DLP is a dose measurement of the total amount of exposure received in a scan series.
Directly proportional to the length of the scan
  Longitudinal anatomy coverage
Must calculate the scan length as well as the CTDI To calculate the average dose received in a series:
DLP = CTDI x Scan Length
Volume
Volume

## Slide 21: Calculating exposure of a scan series
A Helical CT scan was completed over a patients Pelvis measuring 250mm, using a 16-
slice MDCT, with a Pitch of 1.2. Calculate the DLP.
CTDIVol =
Scan Length =
DLP = ?
DLP =
mGy/cm
CTDI | 
Calculates dose of a Helical scan slice | 0.012Gy
Modification to accommodate smaller slice widths | 0.027Gy
Dose index as it relates to number of slices & slice width used | 0.031Gy

## Slide 22: Calculating exposure of a scan series
What is the DLP of a 25 second Helical CT scan of a patients Abdomen, using a 16- slice MDCT, 1.5mm slice thickness, 1second rotation time, and a Pitch of 1.
CTDIVol = 0.012Gy
Scan Length = ? DLP = ?
AC = ? P = 1
Tat = 25 seconds Rotation time = 1 second Number of slices = 16 Slice thickness = 1.5mm
CTDI | 
Calculates dose of a Helical scan slice | 0.012Gy
Modification to accommodate smaller slice widths | 0.027Gy
Dose index as it relates to number of slices & slice width used | 0.031Gy

## Slide 23: III. ED
A measurement that attempts to correlate the amount of dose absorbed by the patients tissues during a CT scan, to the probability of developing biologic effects.
Risk assessment
  Radiosensitivity of organs varies
ED = DLP x k
Compares doses produced during CT scanning to those of natural background exposures

## Slide 24: Calculating Risk for BERT
Calculate ED for a 220mm Abdominal Helical CT scan, that used eight 2mm slices, and a Pitch of 1.
K Factors |  | Weighting Factors | 
Head | 0.0019 | Stomach | 0.12
Neck | 0.0048 | Gonads | 0.20
Chest | 0.012 | Thyroid | 0.05
Abdomen | 0.016 | Bone | 0.01
Pelvis | 0.016 | Lung | 0.12
CTDI | 
Calculates dose of a Helical scan slice | 0.032Gy
Modification to accommodate smaller slice widths | 0.017Gy
Dose index as it relates to number of slices & slice width used | 0.021Gy
CTDIVol =
DLP =
Scan Length =
ED = ?
ED =
Sv

## Slide 25: Things to Take Away
What is dosimetry?
Importance of dosimetry
How are dose measurements obtained?
  Types of dosimeters
  Phantom specifics
CT dose descriptors
  CTDI
  DLP
  ED
Calculations